## Slide 1
Guidelines for open-source shim amplifier setup
Update: December 1, 2018
Jason Stockmann
A. A. Martinos Center for Biomedical Imaging
Massachusetts General Hospital
email questions to: jstockmann@mgh.harvard.edu
MR Physics and Instrumentation Group

## Slide 2
Open-source shim amplifier board circuit board files, schematics, and simulation software available here:
http://rflab.martinos.org/index.php/Current_driver:Current_driver
Bill of materials / parts list for the three circuit boards is also available:
8ch amplifier board
Teensy 3.5 Microcontroller daughter board / fiber optic board
Fiber optic interface board
In the following few slides, we provide suggested parts to use for the remaining components of the amplifier setup, including AC/DC converter power supply, enclosures, cabling, etc.
We also list the parts on the 8ch amplifier boards that don’t come pre-populated on the boards supplied with the open-source kit.

## Slide 3
Subsystem | Part name | Suggested part number | Link | Cost ($ USD) | Notes
Option 1: Linear power supply | Tripp Lite PR60 DC Power Supply 60A 120V AC Input to 13.8 DC Output TAA GSA | PR60 | https://www.alliedelec.com/product/tripp-lite/pr60/70101752/?gclid=Cj0KCQjws7TqBRDgARIsAAHLHP4iqgeWipu5oZA0ZbGap0Tllqs0RZcOtFIktjjRwD2TJ3kcK53_PxUaAqlwEALw_wcB&gclsrc=aw.ds | 340 | Linear supply is the safety bet because you don’t need to worry about filtering out ripple in the switching power supply output
Option 2: Switching Power supply | AC/DC converter, 2000W Single Output Power Supply 12 Volts @ 100 Amps | RSP-2000-12 | https://www.jameco.com/z/RSP-2000-12-MEAN-WELL-2000W-Single-Output-Power-Supply-12-Volts-100-Amps_2152981.html?CID=GOOG&gclid=CjwKCAjwrNjcBRA3EiwAIIOvq4hT5vFfBLhiGqZbjJncPI57JO-Bi-CI3hmqy6vK_Mqv8c29MJvcvRoClx8QAvD_BwE | 375 | Switching supplies can supply higher currents but you need to be careful about the ripple, which changes with output current level and can show up in the image noise floor. The noise floor then can change with the power supply current unless careful filtering is used. Another good option if you only need 50 amps: https://www.jameco.com/z/SE-600-12-MEAN-WELL-AC-to-DC-Power-Supply-Single-Output-12-Volt-50-Amp-600-Watt_374011.html
 | Bud Industries rack-mount Enclosure (13”x17”x3.5”) | 563-NHC-14155 | https://www.mouser.com/ProductDetail/Bud-Industries/NHC-14155?qs=%2fha2pyFaduj9zMZIBX%252bDj2IzVART0WEz1YCyw2YIou8%3d | 109 | Meanwell AC/DC 12V/100A is a tight fit in this enclosure.
 | Simpson DC Ammeter (0-100A) | 02740 | https://www.mouser.com/ProductDetail/Simpson-Electric-Company/02740?qs=sGAEpiMZZMtUpMOmSPzkHy4N5Z8yXGvg | 96 | Can also use slightly smaller 02540 model.
 | Current shunt, Analog Panel Meters SHUNT 100 AMP 50 MV | 06713 | https://www.mouser.com/ProductDetail/Simpson-Electric-Company/06713?qs=sGAEpiMZZMtFJx2zDaGuWCGrWWAP9a5liEaW4QXVCq4%3d |  | 
Linear power supply option | AC/DC CONVERTER 13.8V 828W | TL241-ND | https://www.digikey.com/product-detail/en/tripp-lite/PR60/TL241-ND/251533 |  | A linear supply is an alternative to a switching supply if RF noise is a concern. The linear supply has “clean” output power so a filter panel might not be required. Common mode chokes should still be used on the cables running from the linear supply into the magnet room. If ferrous chokes are used, be careful to keep them outside the magnet room.
Shim amplifiers | OPA549, IC OPAMP POWER 900KHZ TO220-11 | OPA549T-ND | https://www.digikey.com/product-detail/en/texas-instruments/OPA549T/OPA549T-ND/307860 | 23 | Two power op amps needed for each channel (16 per board)
 | Power connector (10-pin MOLEX) | WM21355-ND | https://www.digikey.com/products/en?keywords=WM21355-ND |  | Mating connector for internal enclosure cabling: https://www.digikey.com/product-detail/en/molex-llc/39-01-2100/WM3704-ND/61385
 | Output connector (2-pin MOLEX) | WM21351-ND | https://www.digikey.com/products/en?keywords=WM21351-ND |  | Mating connector for internal enclosure cabling: https://www.digikey.com/product-detail/en/molex-llc/39-01-2020/WM3700-ND/61315
 | Pins for Mini-Fit Jr. connectors | WM2501CT-ND | https://www.digikey.com/product-detail/en/molex-llc/39-00-0038/WM2501CT-ND/467978 |  | Contacts for AWG18-24 for Mini-Fit Jr. MOLEX connectors for internal wiring inside amplifier enclosure
 | Crimp tool for Mini-Fit Jr. MOLEX | 63819-0900 | https://www.newark.com/molex/63819-0900/hand-crimp-tool/dp/97M9440?MER=bn_level5_5NP_EngagementRecSingleItem_1 | 281 | 

## Slide 4
Subsystem | Part name | Suggested part number | Link | Cost ($ USD) | Notes
Cabling | Shim output cable |  |  |  | 
 | Power cable to amplifiers |  |  |  | 
 | Twisted pair cable Option 1 |  |  |  | 

## Slide 5
Subsystem | Part name | Suggested part number | Link | Cost ($ USD) | Notes
ODU connectors. |  |  | https://www.odu-usa.com/fileadmin/redaktion/downloads/downloadcenter/brochures/ODU-MAC_ZERO_Manual_Mating-0518.pdf |  | Helpful contacts at ODU who can help you purchase the correct part numbers: Hauger John John.Hauger@odu-usa.com Lowe Tom Tom.Lowe@odu-usa.com ODU Zero high-density connectors are available with 42-pin (21 channels) and 90-pin (45 channels) options for each connector. See online documentation for more details.
MIL-SPEC connectors | Male box-mount connector | MS3112E20-16P | https://www.mouser.com/ProductDetail/Amphenol-Industrial/MS3112E20-16P?qs=sGAEpiMZZMv%2fye0hRulZR%2fS5ltx8gSydeTjLcqhhYqo%3d | 37 | To use as 12V power input connector for the amplifier board enclosure
 | Circular MIL Spec Connector 16P #16 SKT CONTACTS | MS3116F20-16S | https://www.mouser.com/ProductDetail/Amphenol-Industrial/MS3116F20-16S?qs=sGAEpiMZZMv%2fye0hRulZR%2fS5ltx8gSydZaYg9yxXhSg%3d | 80 | To connect to male box-mount connector on amplifier board enclosure for 12V power input (must make custom cable using this part)
 | Circular MIL Spec Connector 16P Size 20 Pin Plug Strt Strain Relief | PT06E-20-16P-SR | https://www.mouser.com/ProductDetail/Amphenol-Industrial/PT06E-20-16P-SR?qs=%2fha2pyFadugSneYzzMZUg8xUloEPQyyBQexzX6MEOBwLIRvnztMUBQ%3d%3d | 62 | To connect to 12V power supply output box-mount connector (must make custom cable using this part)
 | Female box mount connector | PT06E-20-16P-SR | https://www.mouser.com/ProductDetail/Amphenol-Industrial/PT06E-20-16P-SR?qs=%2fha2pyFadugSneYzzMZUg8xUloEPQyyBQexzX6MEOBwLIRvnztMUBQ%3d%3d |  | 
Cooling System | Lytron Ambient Cooling System (rack-mount version) | MCS20G01BB | https://www.lytrondirect.com/mcs20g01bb-ambient-cooling-system/ | 1400 | 
 | Lytron cold plate (6-pass), 12”x3.8”, ¼” thick | CP15G05 | https://www.lytrondirect.com/cp15g05-tubed-cold-plate/cp15g05/ | 167 | Carefully drill holes in this cold plate for mounting OPA549 op amps. Put thermal grease between op amps and aluminum cold plate.

## Slide 6
Subsystem | Part name | Suggested part number | Link | Cost ($ USD) | Notes
Plumbing (Swage-Lok) | Brass hose tube adapter 1/4 | B-4-HC-A-401 |  |  | 
 | Vinyl tubing | LT-8-10 |  |  | Vinyl Tubing, 1/2 in. ID - 5/8
 | Brass Reducer, 1/4 in. OD | B-400-R-8 |  |  | 
 | Brass Hose Tube Adapter, 1/2 i | B-8-HC-A-811 |  |  | 
 | Brass Quick Connect Bulkhead B | B-QC8-B1-810 |  |  | 
 | Brass Quick Connect Stem, 1/2 | B-QC8-D-810 |  |  | 
Filter (to be continued…) | Schaffner 200 amp filter capacitor | FN 7563-200-M10 | https://www.digikey.com/products/en?FV=ffec4bdd |  | Be sure to filter both the input power line and return path. Use filter capacitor that is mounted to the faraday shield of the magnet room.
 | Schaffner 100 amp filter capacitor | FN 7563-100-M8 |  |  | 

## Slide 7
Console room
Shim amplifiers
Magnet room
Arduino microcontroller
Fiber optic transmitter/receiver
Fiber cable
Equipment room
Switch-mode Power supply (12V, 80A)
Filter
TTL trigger from scanner
ΔB0/Rx array coil in magnet bore
Cabling for 31 shim channels
SPI Bus
USB port
Faraday cage
Fiber optic transmitter/receiver
RF waveguide
Optional pump for water cooling
SPI Bus
Tygon tubing
Figure. Experimental setup of dynamic B0 shimming hardware. ΔB0 field maps are processed off-line on a laptop computer and optimal shim currents for each slice are computed. The shim settings are sent to an Arduino microcontroller which then sends SPI bus commands to the digital-to-analog converters on the amplifier boards via fiber optic cable. The amplifiers are placed inside the magnet room to minimize the number of feedthrough filters required. Only the 12V power rail and ground for the amplifiers required filtering. The amplifier outputs are supplied via twisted pair cables to the ΔB0/Rx array coil inside the magnet bore. For slice-optimized, dynamic shimming, TTL pulses supplied by the scanner trigger the microcontroller to update the shim currents 500us before the beginning of each TR. Amplifier output currents are monitored and sent back to the microcontroller via an analog-to-digital converter.

## Slide 8
8ch amplifier board
8ch fiber optic interface board
Arduino microcontroller daughter board
Large order of populated amplifier boards received in Oct. 2018. Only through-hole components remain to be added (power amplifiers and connectors).
Completed amplifier board with heat sink

## Slide 10
(Rev. C, V5) Design: Linear, current feedback, lead-compensated push-pull
Provides ~30μs rise time for driving 10μH shim coil load

## Slide 11
Enclosure/chassis photos

## Slide 12
Filter box on scanner filter panel

## Slide 13
Amplifier chassis photos
Fiber board is mounted vertically to allow short ethernet cable runs to the amplifier boards, which are horizontal

## Slide 14
Use brass rods to support boards on heat sinks
Use non-magnetic fasteners!

## Slide 16
Fiber optic connections
Make sure to add some sort of strain relief for the fiber optic cables

## Slide 17
+12V Power connection

## Slide 18
Steps for installing OPA-549 to avoid mechanical stress on solder joints:
Drill holes in heat sink
Set up PCB on brass threaded rod with nuts to hold it in place. Remember to use thermal compound between op amp and heat sink
Temporarily attach OPA-549 pair to heat sink with screw
Solder pins to PCB
Do not solder the op amps first, and then install the bolts second. This can damage the pins.

## Slide 20
Power supply enclosure

## Slide 21
12V, 100A power supply enclosure with ammeter

## Slide 22
Inside of power supply
Make sure to use fuse on the AC input for safety!!!
Outlet includes fuse

## Slide 23
Cabling considerations

## Slide 24
MIL-SPEC Amphenol 16-pin connector for supplying 12V power to shim amplifiers (through filter panel). Note that these connectors have a ferrous washer to make the bayonet connection mechanism work. Minor issue inside the scanner room.
Suggested pinout:
Pins A-E: +12V
Pins M, S, R, P, N: Ground/return path
Pins L, K, J, H, G, F: Reserve/unused

## Slide 25
ODU cable for shim channels. Adjacent pins are connected to one twisted pair inside cable. Each shim channel requires a separate return path.
ODU does not provide enclosure for the receptacle so you need to build your own from a box and a gray Hubble cable gland for strain relief

## Slide 27
Notes on ODU cables:
Three blocks each with 14 pins = 42 pins total for 21 shim channels.
Each shim channel requires a separate return path. DO NOT CONNECT SHIM CHANNEL RETURN PATHS INTO A COMMON GROUND INSIDE THE SHIM COIL OR INSIDE THE SHIM AMPLIFIER ENCLOSURE.
NOTE: 90-pin versions of this connector are also available from ODU. You can change the inserts to use thinner insulating bodies that have higher pin density. The ODU connector blocks are easily changeable.

## Slide 28
Teensy Arduino control software

## Slide 29
Teensy Arduino control software: runCRev
In Arduino software, choose: Tools  Board  Teensy 3.5
Then Tools  Port  Teensy port
This is where you set the output current in the runCRev file:
Number of “slices” to update through (for optional dynamic shimming)
Number of shim channels (columns). Can go up to 64ch.
Board 0
Array of currents in each channel (columns) for each slice/trigger (rows)
Board 1
Board 2
Board 3

## Slide 30
(1) Click here to compile with Teensy connected
(2) Click here open serial port dialog box

## Slide 31
To enable triggering for dynamic shim updating, uncomment this line. Must connect via fiber optic or BNC to scanner trigger output (“SYNC” on Siemens).

## Slide 32
Serial port dialog box.
These four single-characters commands are how you control the Teensy:
“I” to read out currents on each channel
”C” performs auto-calibration to zero currents on each channel
“Z” zeros all channels
“M” loads the first set of shims and ensures that Arduino is in triggered mode
Outputs:
First column is board and channel addresses
Second column is proportional to the input voltage on each channel
Third column is proportional to ADC output which reads currents sense resistor
Don’t worry too much about the absolute value of these numbers. Just make sure all channels pass auto-calibration step.
Watch out for channels that fail the auto-calibration! This either means (a) the channel is disconnected, (b) the channel is broken, or (c) the channel is deliberately unused

## Slide 33: Guidelines for dynamic slice-optimized shim updating

## Slide 34
Hz
Delta B0 maps
Hz
Fat sat
ΔB0 map with MC shim field
Baseline ΔB0 map
Axial
Sagittal
Figure. Brain ΔB0 field maps with and without an applied multicoil shim field offset that is optimized for the 1cm slab bounded in the purple box. While the multicoil field offset improves ΔB0 homogeneity within the target slice (top right), it creates large field offsets elsewhere in the brain. In parts of the frontal lobes, the multicoil field shifts the water into the saturation band of the fat saturation pulse (-350 Hz lower than water). This leads to spin history effects and reduced contrast when the superior slices of the brain are acquired during subsequent TRs.
ΔB0

## Slide 35
Anatomic
Static shim
Naïve dyn. shim
Optimal dyn. shim
Figure. EPI slices showing loss of contrast due to spin history effects caused by updating the slice-optimized multicoil shim currents before the fat saturation pulse during each TR. The problem is solved by updating to slice-optimized shim currents in between the fat saturation and water excitation pulses, and playing a whole-head global shim during fat saturation. Parameters: Gradient echo EPI, 2x2mm in-plane with 4mm slice, TE=32ms.

## Slide 36
Excitation
Fat sat
RF Tx
Slice select
Readout
Phase encode
One shim channel
Naïve trigger
Optimal trigger
Figure. Gradient-echo echo planar imaging sequence showing the “naïve” and optimal triggering schemes for slice-to-slice multi-coil shim updating. Updating to a slice-optimal shim before the fat saturation pulse can lead to severe artifacts due to spin history effects outside the slice being shimmed. Setting the multicoil shims to zero or to a whole-head shim during the fat saturation pulse, and then updating to the slice-optimal
(whole-head shim)
(slice-optimal shim)

## Slide 37: Other artifacts
Subtle artifacts can occur during EPI readouts as the amplifier boards try to compensate for the strong readout gradient slew that induces back-EMF in the shim coils inside the bore. See following slide.
We observed horizontal lines along the readout direction (not PE direction) that seem to come from a k-space disturbance that is periodic in time during even and odd line readouts in EPI
The disturbance may be due to the amplifier ground varying relative to the scanner RF ground, causing a disturbance that is detected by the RF preamps.
The artifact disappears if the two grounds are connected through a capacitor (47nF). Do not directly connect the grounds; only capacitive coupling is permitted. Check with your scanner vendor support engineering or site MR physicist for approval before capacitively-coupling the grounds. We have not had any issues at our site.
This artifact has only been observed on integrated B0/Rx coils where the same loop is used for shimming and RF receive. It may not be a problem for shim-only loops.

## Slide 38
Separate grounds
Connected grounds
GRE-EPI volumes
Standard window
10x window
0
1000
0
100
Image mag. (a.u.)
Phase-encode direction
Horizontal line artifact!

## Slide 39: Details on board testing
Reproduced from Arango N, ISMRM 2016

## Slide 40: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
15 cm
from isocenter
Position test loop 15 cm from isocenter and oriented for maximum back-EMF caused by EPI readout (z) gradient slewing to maximize flux
Notes: READOUT ALONG Z SAY THAT THING creating a large amo

## Slide 41: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
Position test loop 15 cm from isocenter and oriented for maximum back-EMF caused by EPI readout (z) gradient slewing to maximize flux
Gradient slew disturbance

## Slide 42: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
Shim current during gradient switching
Closed loop controller rejects gradient slew

## Slide 43: Results: Dynamic disturbance rejection
Sequence parameters: EPI GRE sagittal, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
3T scanner bore
Shim current during gradient switching
Closed loop controller rejects gradient slew

## Slide 44: Results: Shim current can be set 500 𝜇s before image TR
Sequence parameters: EPI GRE axial, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
+2A
0A
-2A
0A
Single-turn test loop
Phantom
3T scanner bore
Signal ROI
Noise ROI
Test: Update shims to +/-2A starting 500us before beginning of TR in EPI-GRE sequence
Current driver
EPI magnitude images (20 slices)
a.u.
shim current OFF
SNR = 48.0
shim current ON
SNR = 48.5
Notes: Shows correct distortion in the phase encode direction

## Slide 45: Results: No excess RF noise caused by dynamic shimming
Sequence parameters: EPI GRE axial, 3mm slices ascending, 1.6mm in-plane, matrix 128x128x20, 1302 Hz/pix, 0.94 echo spacing, 6/8 partial Fourier, TE=37 ms, TR=9000 ms
Current driver
Single-turn test loop
Phantom
3T scanner bore
Test: Update shims to +/-2A starting 500us before beginning of TR in EPI-GRE sequence
→ No RF “zipper” artifacts or structured noise
a.u.
+2A
0A
-2A
0A
EPI magnitude images (20 slices)
shim current OFF
SNR = 48.0
shim current ON
SNR = 48.5

## Slide 46: Frequently asked questions
========================================
Is the unit set up for bipolar or unipolar operation (on the power supplies)?
The shim array amplifiers are designed to be powered with one single ended power supply and provide bipolar outputs through the use of the bridge configuration. In practice we don’t recommend going above 24V without checking the capacitor voltage ratings in the bill-of-materials. Also, higher voltages will tend to dissipate more heat in the power op amps (assuming a relatively small load resistance).
2) What type of connectors are used to get the rather heavy current on and off the board?
Each channel is designed for +-3.5A maximum so the molex minifit 90 degree connectors used with 9A ratings are more than sufficient. If channels are used in a parallel configuration to share current and increase maximum current per channel, outputs are connected off-board so the same connectors are used.
3) Was the bridge configuration chosen simply because you need to drive rather high resistance loads?
The bridge configuration makes single supply input possible eliminating the need for a virtual ground reference. This design choice also makes heat-sinking simpler as the heat sink tabs on the OPA549 are electrically connected to their V- pin. By running a bridge with a single supply, both OPA549s used per channel share V- and may be connected to the same heat sink without the potential for costly error for users building this themselves with the use of insulating heat-sink pads
4) We have +/- 15 volt outputs from our coil drive. Is it possible to scale this and go in after your DAC? I ask this not because of electrical issues but rather access to pcb traces can be a problem.
Yes it is possible to not use the DAC and instead drive each channel separately with an appropriately scaled and shifted analog input. Zero ohm resistors separate DAC outputs from each amplifier input and plated through holes allow users to feed in their own analog input. The input cannot be bipolar due to the single supply operation of the device. With the DAC in place, the inputs are differential with respect to a 1.25V DAC reference.
5) Can you recommend a fan cooled radiator and pump to complete the cooling loop?
The only constraint here is that the spacing between the OPA549 op amps is set for a particular cooling plate (1/4” nominal thickness). It’s a LytronCP15G05 6-pass cooling plate. We have a Lytron MCS20G01BB water cooling system which can dissipate up to ~1kW, but in practice, to simplify our setup we haven’t been using it for our shimming experiments. Whether the water cooling will be necessary depends on how aggressively you set the output currents and how long you sustain them. However at the end of some of our sessions, the cooling plates are definitely quite toasty.
So for initial testing, an aluminum block should be adequate, but depending on your duty cycle & output currents you may need to consider water cooling for the long-run.
6) Is the published schematic a simplified one? I expected protection diodes and snubbers.
I’m sorry about this oversight. The published schematic is not the latest version though the latest version does not include input protection diodes or snubbers. The quiescent current draw is not so large that we see appreciable voltage kicks on power-on. If you mean on the outputs, it is important to note the behavior of the lead compensated feedback networks which enable stable current-mode drive of loads over the inductance range for which this amplifier is appropriate. Changes to the lead compensation network enable stable control from 1uH to 900uH with bandwidths up to 20kHz. Protection diodes would be appropriate in a switch-mode design to protect sensitive MOS devices from over voltage or for high vibration environments in which intermittent connections while driving large current through high inductance loads is expected.
7) Do you build the device and sell or share schematics for us to build if we choose?
If you choose to do so, you are completely welcome to build this device at no cost using our open-source design files.
We do have a limited number of populated 8ch amplifier boards and we’d be happy to provide you with one or more boards at cost, if it would be helpful. The fiber boards and microcontroller daughter board still need to be populated, but this is more straightforward than populating the amplifier board (which has many more small components including fine-pitch ICs that are hard to solder by hand). Let me know what your project involves and I could help put together a kit of parts to mail to you.
8) Is there any way to correct for the op amps offset voltage?
When operated with the DAC and ADC, the channel to channel gain and DC offset variation are calibrated automatically. Opamp DC offset is a function of device temperature so single set-point trimming can be ineffective for power applications. In our shimming work the calibration takes less than 100ms and is run before beginning each scan (perhaps overkill but not burdensome).
We do have a second amplifier topology which includes an integrator to zero out offsets:
http://tabletop.martinos.org/index.php/Hardware:GPA
This single-channel design is somewhat similar to what is used in Techron gradient amplifiers (though stripped down to the bare bones) . However, this circuit amplifier board is still very “basic” and does not have all the features of the 8ch shim amplifier board (which itself evolved out of a motor controller circuit on an L272 datasheet). The microcontroller software for the 8ch board also has more features than our software for the single-channel amplifier.
Our recommendation is to go with the 8ch Rev. C board and use the autocalibration feature in the control software. However if you’d like to try our older amplifier design, we are happy to provide PCBs for that design as well.
Alexandru,
The cold plate we use is ¼” thick and fits properly on the Rev. C-Ver5 board: https://www.lytrondirect.com/cp15g05-tubed-cold-plate/cp15g05/?b=549 Are you reading the 3.8” dimension maybe? That refers to the height of the cold plate. The copper pipe is ¼” thick. This cold plate should work.
We use the standard Siemens fiber optic cables. They come in bundles of 5 or 10 fibers usually in an orange jacket. If you’re lucky, your Siemens site engineer might be able to find some spares. Otherwise I could ask Thomas if we might be able to mail you some. Otherwise you could call Mouser or Digikey and ask for Avago-compatible fiber optic cables. Otherwise you could gang a few individual fibers together yourself with heat shrink to make a cable, though admittedly this isn’t as nice to work with. For example, would 5 meters be long enough for your setup? If not, you can try using fiber couplers to extend the reach.
https://www.digikey.com/products/en?mpart=HFBR-RNS005Z&v=516
https://www.mouser.com/ProductDetail/Broadcom-Avago/HFBR-4505Z?qs=jT9z6tsiFNkQ5IvqQth7%252BA%3D%3D
I use a USB-to-micro-USB to connect my laptop to the Arduino directly. I put the Arduino in a small box with a micro-USB extender cable that runs from the Arduino itself to a micro-USB mounted to the enclosure panel. The box sits on the desk next to my computer. The fiber then runs into the scanner room to the fiber interface board, which lives with the amplifiers. This works well as long as the fiber is long enough and as long as you don’t use the DACs or ADCs during signal readout (this can cause RF noise). See attached diagram for the layout.
Right now I’m not using U013 or U014. The purpose of U012 is to receive triggers from the MRI scanner to tell the Arduino to update the shims to the next slice setting (or for other field control applications). The BNC connector is for general purpose use. I use it for triggers fMRI experiments because we use the scanner’s SYNC output trigger pulse on U012 to do dynamic shimming. Normally this fiber goes to a pulse stretcher with BNC output that goes to our fMRI stimulus computer (Julien will know how this works). So our Arduino BNC output just provides a stretched trigger pulse to tell the fMRI experiment when to start. I don’t know if this will matter for your applications, but I can show you how to enable it in the Arduino software.
I like to put the Arduino board in an enclosure like this one below. Just check to make sure the board will fit in the rails inside the enclosure. These are really nice because you can remove the panels to drill or machine the holes for the connectors.
https://www.digikey.com/product-detail/en/hammond-manufacturing/1455L1202/HM895-ND/1090707
https://www.digikey.com/product-detail/en/hammond-manufacturing/1455N1202/HM975-ND/460175
Here is a micro-USB extender that is panel-mount (it’s a little long, but you can loop it around inside the enclosure):
https://www.digikey.com/product-detail/en/adafruit-industries-llc/3258/1528-1786-ND/6238006
X003 supplies +12V power to the fiber interface board (then there are 5V and 3.3V regulators on the fiber board to step the 12V down as needed). You are correct, this uses the Molex 2-pin “mini-fit jr.” PCB-mount connector. This is the same voltage that the amplifier board runs on. You can also use a little higher voltage if you want. If you go above 16V you need to make sure all the ceramic 10uF and 100uF capacitors are rated to go above 16V on the main amplifier board (I think they go to 25V…).
--- Jason
From: Alexandru Foias <alexandrufoias@gmail.com> Date: Thursday, March 7, 2019 at 11:05 AM To: Jason Stockmann <bauhaus2@gmail.com> Cc: <ACDC_SPINE_7T@googlegroups.com> Subject: shim boards questions
Hi Jason,
I have some questions regarding the shim boards.
Shim board: In the PPT you recommend the heatsink CP15G05 TUBED COLD PLATE , which is 3/8" wide. Is it compatible to rev C ver5 board ?
Connection interface board - teensy board. What type of fiber optics do you recommend for the link between interface board - teensy board. I saw that you are using a hand made mutlifilar FO cable. Is it better to use that one or a single FO cable ?
Teensy board: How do you connect the teensy to the PC ? We would like to extend the serial port of teensy using fiber optics (https://www.pjrc.com/teensy/td_uart.html). There are 3 receiver FO (U012, U013, U014) and a BNC connector (U002) on the Teensy board. Could you please tell us what is their purpose ?
Interface board. The part with the footprint X003 is not specified in the BOM. It seems to be similar to header in the shim board MOLEX 0026013114. Could you please confirm if that's the case? What voltage should we use to feed the interface board ?
Thanks,
Alexandru
--
Alexandru Foias MScA
Research Associate
Department of Electrical Engineering, Polytechnique Montréal
2900 Edouard-Montpetit Bld, room L-5626
Montreal, QC, H3T1J4, Canada
Phone: +1 (514) 340-4711 (office: 7546)
E-mail: alexandru.foias@polymtl.ca
Web: www.neuro.polymtl.ca